## Slide 1
38655 BMED-2300-02
Lecture 14: CT Scanner
Ge Wang, PhD
Biomedical Imaging Center
CBIS/BME, RPI
wangg6@rpi.edu
March 9, 2018

## Slide 2
Tue | Topic | Fri | Topic
1/16 | Introduction | 1/19 | MatLab I (Basics)
1/23 | System | 1/26 | Convolution
1/30 | Fourier Series | 2/02 | Fourier Transform
2/06 | Signal Processing | 2/09 | Discrete FT & FFT
2/13 | MatLab II (Homework) | 2/16 | Network
2/20 | No Class | 2/23 | Exam I
2/27 | Quality & Performance | 3/02 | X-ray & Radiography
3/06 | CT Reconstruction | 3/09 | CT Scanner
3/20 | MatLab III (CT) | 3/23 | Nuclear Physics
3/27 | PET & SPECT | 3/30 | MRI I
4/03 | Exam II | 4/06 | MRI II
4/10 | MRI III | 4/13 | Ultrasound I
4/17 | Ultrasound II | 4/20 | Optical Imaging
4/24 | Machine Learning | 4/27 | Exam III
BB Schedule for S18
Office Hour: Ge Tue & Fri 3-4 @ CBIS 3209 | wangg6@rpi.edu
Kathleen Mon 4-5 & Thurs 4-5 @ JEC 7045 | chens18@rpi.edu

## Slide 3: CT Scanner
Data Truncation
CT Scanning
CT Generations
Spiral/Helical CT
Interior Tomography
CT Image Artifacts
Image Quality
Image Artifacts
X-ray Radiation Dose
Dose Measures
Low-dose CT

## Slide 4: Sufficient Condition: 2D
At least a point on a source orbit
on any line in an object

## Slide 5: Sufficient Condition: 3D
At least a point on a source orbit
on any plane in an object

## Slide 6: Classic Scheme / Central Dogma
A whole cross-section or an entire object must be completely covered by an x-ray beam (parallel-beam, fan-beam, or cone-beam)
Resultant data are transformed to the Fourier space that must be fully sampled for image reconstruction over the section or the object

## Slide 7: CT Scanner
Data Truncation
CT Scanning
CT Generations
Spiral/Helical CT
Interior Tomography
CT Image Artifacts
Image Quality
Image Artifacts
X-ray Radiation Dose
Dose Measures
Low-dose CT

## Slide 8: First Four Generations

## Slide 9: Electron-beam CT

## Slide 10: Seven Generations
Generation | Source | Source collimation | Detector | Detector collimation | Source-detector movement | Advantages | Disadvantages
1G | Single x-ray tube | Pencil beam | single | none | Move linearly and rotate in unison | Scattered energy is undetected | slow
2G | Single x-ray tube | Fan beam, enough to cover FOV | multiple | Collimated to source direction | Move linearly and rotate in unison | Faster than 1G | Lower efficiency and larger noise because of the collimation in detectors
3G | Single x-ray tube | Fan beam, enough to cover FOV | Many | Collimated to source direction | Rotate in synchrony | Faster than 2G, continuous rotation using a slip ring | More expensive than 2G, low efficiency
4G | Single x-ray tube | Fan beam, covers FOV | Stationary ring of detectors | Cannot collimate detectors | Detectors are fixed, source rotates | Higher efficiency than 3G | High scattering since detectors are not collimated
5G (EBTC) | Many tungsten anodes in single large tube | Fan beam | Stationary ring of detectors | Cannot collimate detectors | No moving parts | Extremely fast, capable of stop-action imaging of beating heart | High-cost, difficult to calibrate
6G (Single-slice Spiral CT) | 3G/4G | 3G/4G | 3G/4G | 3G/4G | 3G/4G plus linear patient table motion | Fast 3D images | A bit more expensive
7G (Multi-slice Spiral CT) | Single x-ray tube | Cone beam | Multiple arrays of detectors | Collimated to source direction | 3G/4G/6G motion | Fast 3D images | expensive

## Slide 11: Scanning Modes
Projection data can be truncated transversely or longitudinally, which makes it difficult formulating the Fourier domain sampling patterns

## Slide 12: Spiral Single-slice CT

## Slide 13: Full-scan Interpolation
x
y
z
Ray A
Full-scan projection data can be interpolated into a complete set of projection data on an imaging plane such as the x-y plane. The key is to find neighboring rays of the same orientation.

## Slide 14: Half-scan Interpolation
x
y
z
Ray A
Half-scan projection data can be interpolated into a complete set of projection data on an imaging plane such as the x-y plane. The key is to find neighboring rays in the opposite orientations.

## Slide 15: Retrospective Reconstruction
Scanning Loci
Slices
Scanning Locus
Slices
Direct
Reconstruction
Retrospective
Reconstruction
Incremental (Left) vs Spiral (Right) Scans Define Imaging Planes Differently. The Former Specifies Imaging Planes Physically/Prospectively, while the Latter Does so Computationally/Retrospectively.

## Slide 16: Superior Detectability
Direct
Reconstruction
Retrospective
Reconstruction
Retrospective Reconstruction Gives Better Lesion Detectability If There Are Sufficiently Many Slices Reconstructed!

## Slide 17: Theoretical Conclusion
“For a given X-ray dose, helical CT allows substantially better longitudinal resolution than conventional CT due to its inherent retrospective reconstruction capability.”
        Wang and Vannier
        Medical Physics 21:429-433, 1994

## Slide 18: Spiral Multi-slice/Cone-beam CT
Wang, G, Lin, TH, Cheng PC, Shinozaki DM, Kim HG: Scanning cone-beam reconstruction algorithms for x-ray microtomography. Proc. SPIE 1556:99-112, July 1991
Wang G, Lin TH, Cheng PC, Shinozaki DM: A general cone-beam reconstruction algorithm. IEEE Trans. on Med. Imaging 12:486-496, 1993

## Slide 20: Citation Counts
To solve the long-object problem, a first level of improvement with respect to the 2D FBP algorithms was obtained by backprojecting the data in 3D, along the actual measurement rays. The prototype of this approach is the algorithm of Wang et al.
Defrise, Noo, Kudo: A solution to the long-object problem in helical cone-beam tomography. Phys. Med. Biol. 45:623-643, 2000
Many advances in CB reconstruction have been made recently thanks to the quest for an attractive reconstruction method in helical CB tomography.
Pack, Noo, Clackdoyle: Cone-beam reconstruction using the backprojection of locally filtered projections. IEEE Trans. Medical Imaging 24:1-16, 2005
100M Multi-slice/Cone-beam CT Scans Annually

## Slide 21: Longitudinal Data Truncation

## Slide 22: Exact Spiral Cone-beam CT
Detector
Plate
Source
Object
Pi-Line
Katsevich A: An improved exact filtered backprojection algorithm for spiral computed tomography. Advances in Applied Math. 32: 681-697, 2004

## Slide 23: Transverse Data Truncation
X-rays

## Slide 24: CT vs Interior Tomography
The Generic Interior Problem Has No Unique Solution in an Unconstrained Functional Space. An Additional Assumption is Needed:
(1) There Is a Small Sub-region Over Which CT Number is Known; or
(2) The Region of Interest (ROI) is Piecewise Polynomial, or
… …
All Data vs Partial Data (Truncated Both Sides)
Global vs Local Reconstruction
Whole Sale vs Retail
Door to New Opportunities
http://iopscience.iop.org/article/10.1088/0031-9155/58/16/R161/pdf

## Slide 25: Interior Tomography
t
Sinogram
Detector Array
Line Integrals
Reconstruction
X-rays
Object
t
We Revisited the Long-Standing Interior Problem and Developed the Theory of Interior Tomography to Enable Exact Image Reconstruction over an Interior Region of Interest from Truncated Projection Data.

## Slide 26: Less = Faster
Wang G, Yu HY, Ye YB: A scheme for multi-source interior tomography. Med. Phys. 36:3575-3581, 2009
Wang G: The meaning of interior tomography. Proc. of ICASSP, 5764-5767, 2011
Ritman EL, Kinsey JH, Robb RA, Gilbert BK, Harris LD, Wood EH: Three-dimensional image of the heart, lungs, and circulatory. Science 210:273-280, 1980
Less =
Deeper understanding
Larger object
Less radiation dose
Finer image resolution
Wider dynamic range
… …
Faster imaging speed

## Slide 27: CT Scanner Architecture

## Slide 28: Typical Diagram

## Slide 30
Background from Best Conventional CT
ROI Info from Spectral Interior Tomography
In-vivo Spectral Micro-CT Enabled with
– Parallel Localized Data Acquisition
– Hybrid Advanced Reconstruction
Cardiac CT Architecture

## Slide 31: CT Scanner
Data Truncation
CT Scanning
CT Generations
Spiral/Helical CT
Interior Tomography
CT Image Artifacts
Image Quality
Image Artifacts
X-ray Radiation Dose
Dose Measures
Low-dose CT

## Slide 32: Data Sampling
~1K Rays per View, ~1K Views
Cross-sectional Image: 512x512 Pixels, ~4K Levels

## Slide 33: CT Number
CT Image Quantized to CT Number in Hounsfield Unit:

## Slide 34: Examples

## Slide 35: Image Resolution
High-contrast
  - Distinguish small objects
  adjacent to each other
Low-contrast
  - Differentiate regions
    similar in gray-level
Temporal resolution
  - Capture time-varying structures

## Slide 36: Image Noise
Due to Quantum and Other Random Effects
Depending on
  Collimation
  Tube Voltage
  mAs Product
  Patient Size
  HVL for Water 3.6 cm
  Voxel Size
  Algorithm

## Slide 37: Low-contrast Resolution
Depending on
  Object Size
  Object Contrast
  System MTF
  System SSP
  Image Noise

## Slide 38: Temporal Resolution
Spiral Scanning
Through-plane motion artifacts are
suppressed by single-breath-hold spiral scanning
Depending on
  Scanning speed
  Reconstruction method
Conventional Scanning
(From JA Brink at Yale U)

## Slide 39: Scatter Artifacts
Randomly Scattered X-ray Photons Can Reach Detectors, Forming a Cloudy Background to Compromise Contrast Resolution. Anti-scatter Grid and/or Software Correction are Needed.

## Slide 40: Beam-hardening
X-ray path
KeV
KeV
KeV
Longer X-ray Path/Stronger X-ray Attenuation Decreases Low Energy Components More Than High Energy Counterparts, Making the X-ray Spectrum Harder (Higher Energy in Greater Portion) after the Attenuation Process.

## Slide 41: Beam-hardening Artifacts
(From J Hsieh at GE)
Without correction
With correction

## Slide 42: Volume Averaging
(Blurred data from GH Esselman at Wash U)
Blurred
Deblurred
Volume averaging
Finite Focal Spot, Finite Detector Aperture, Finite Pixel Size, and Many Other Factors Blur Structural Details.

## Slide 43: Metal Artifacts
(From DD Robertson at Wash U)
Metal Parts Cause Strong Beam-hardening and Severe Photon Starvation (Few Photons Arriving at Detectors, Leading to a High Level of Statistical Noise).

## Slide 44: Motion Artifacts
q=0o
q=90o
Filtered
Backprojection
Iterative
Reconstruction with
Known Motion Pattern
Time-varying Phantom
Motion Artifacts Come from Model Mismatch. An Underlying Structure Keeps Changing While Data Are Collected. The Motion Pattern Can be Taken into Account for Accurate Reconstruction!

## Slide 45: CT Scanner
Data Truncation
CT Scanning
CT Generations
Spiral/Helical CT
Interior Tomography
CT Image Artifacts
Image Quality
Image Artifacts
X-ray Radiation Dose
Dose Measures
Low-dose CT

## Slide 48: Natural Reference
http://nuclearsafety.gc.ca/images/radiation-information/graph-average-annual-effective-dose-natural-sources.gif

## Slide 50: Radiation Dose
Absorbed Dose - Radiation Energy
Effective Dose – Organ Sensitivity
Dose Equivalent – Beam Quality

## Slide 51: Radiation Profile: Single Scan
Radiation spreads outside
the designated slice due to scattering
Ideal profile
Real profile

## Slide 52: CT Dose Index (CTDI)
CTDI: CT dose index
T: slice thickness
D(z): local dose
z: longitudinal coordinate
T

## Slide 53: Radiation Profile: Multiple Scans
Radiation Dose from Multiple Scans Are
Accumulated in the Central Slice
Real profile

## Slide 54: Multiple Scan Average Dose (MSAD)
MSAD: multiple scan average dose
I: inter-slice distance
N: Number of scans
I

## Slide 55: Dose Measurement
Cylindrical Phantoms of 16 cm & 32 cm
Pencil Ionization Chamber
Dosimeter
16 or 32 cm

## Slide 56: MSAD Estimation
MSAD is
  Directly Proportional to mA
  Directly Proportional to Scan Time
  Increases with kVp
  Increases Slightly as Slice Thickness Decreases
  Similar at the Iso-center and Near Surface for Head
  Significantly Less at the Iso-center than
    Near Surface for Body

## Slide 57: Low-dose CT Roadmap (2012)

## Slide 59
BB14 Homework
Green Book
1.14, 1.15 and 1.18
Optional:
As an Art_X Project, Please Design Portable CT Scanner in an Auto-driving Car, or a CT Scanner for a Novel Application.
Due Date: Same (1 Working Week Later)
Enjoy Spring Break!

## Slide 60: Two Examples
https://arxiv.org/ftp/arxiv/papers/1312/1312.6046.pdf
http://content.iospress.com/articles/journal-of-x-ray-science-and-technology/xst00453