## Slide 1
Nuclear Imaging
Ge Wang
AI-based X-ray Imaging System (AXIS) Lab
Rensselaer Polytechnic Institute, Troy, New York, USA
November 9, 2021

## Slide 2: MI Schedule for F21
Tue | Topic | Fri | Topic
08/31 | Introduction | 09/03 | Neuron, Network & Backpropagation
09/07 | No Class | 09/10 | Hands-on 1: Python, Colab, & TensorF
09/14 | Architectures & Training | 09/19 | Hands-on 2: MNIST Classification
09/21 | Image Quality I | 09/24 | CT Physics
09/28 | CT Reconstruction | 10/01 | Deep CT Reconstruction
10/05 | Hands-on 3: CT Networks | 10/08 | More CT Contents
10/12 | MRI Physics | 10/15 | Spin-Echo
10/19 | K-Space Theorem | 10/22 | More MRI Contents
10/26 | Mid Exam | 10/29 | Deep MRI Reconstruction
11/02 | Hands-on 4: MRI Networks | 11/05 | Nuclear Physics
11/09 | Nuclear Imaging | 11/12 | US Imaging
11/16 | Optical Imaging | 11/19 | Multimodality Imaging
11/23 | Hands-on 5: Image Conversion | 11/26 | Thanksgiving
11/30 | Image Quality II | 12/03 | Stability, Interpretability, & Others
12/07 | Overview | 12/10 | Final Exam
Office Hour: Teaching Day 5-6pm: https://rensselaer.webex.com/meet/wangg6

## Slide 3: Polarization & Manipulation

## Slide 4: Existing Polarizable Tracers
Characteristics of some radioisotopes that are the best known candidates for SET at this time. However, these would not be practical for clinical imaging. |  |  |  | 
Isotope | Half-life | Spin number | γ-ray Emission (keV) | γ-ray Branching Ratio
131mXe | 12 days | 11/2 | 164 | 2%
127mXe | 69 sec | 9/2 | 124.8-172.5 | 38%
79mKr | 50 sec | 7/2 | 130 | 27%

## Slide 5: Preferred Properties of Tracers
Polarity: 60-90%
Half-life of Radioactivity: Minutes to Hours
Half-life of Polarity: Minutes to Hours
Biocompatibility: With Possible Surface Modifier
Cost-effectiveness: Comparable to Other Tracers
… …

## Slide 6: SET Idea

## Slide 7: Spin Around

## Slide 8: SET Data Model
At location r
m1 and m2: paired data
Φ: γ-ray flux
µ: attenuation coefficient
Linear integral equation permits inverse solution of
Attenuation coefficients
Flux escaping each side of activated location proportional to concentration
m1
m2
r

## Slide 9: Quiz
If we turn on each pixel at a time and measure data on two sides, can we solve all µ and λ values uniquely?
I: flux
µ: attenuation coefficient
λ: concentration

## Slide 10: Full-voxel Activation
Turned On

## Slide 11: Half-voxel Activation

## Slide 12: Outline
Data Modeling
Statistical Reconstruction
Heuristics

## Slide 13
  Interaction of g photons with matter E ~ 60 - 600 keV (e.g. E 99mTc = 140 keV)
  attenuation is due to Compton scatter and photo-electric absorption
    single photon emission:
    (m = linear attenuation coefficient)
    example: median of penetration depth for 99mTc (140 keV) = 4.5 cm
    photon pair emission (positron emission):
s
s=a
s=d
s
s=a
s=d2
g
g
s=d1
the attenuation is identical for each point along the projection line

## Slide 14
Fourier reconstruction and filtered backprojection (FBP)
  Detected number of photons
  SPECT:
  PET:
  Radon transform:
    Not directly applicable due to attenuation !
  Solution: transmission scan -> m -> FBP for PET
x
y
r
q
f(x,y)
s

## Slide 15: Neat Data Models for an Impulse
s
s=a
s
s=a
s=d2
g
g
s=d1
s=d2
s=d1

## Slide 16: SPECT Model More Complicated
s
s=a
s=d2
s=d1
s=b

## Slide 17: Outline
Data Modeling
Statistical Reconstruction
Heuristics

## Slide 18
Iterative reconstruction (also for 3D)
  Bayesian approach:
  Given L : reconstructed image Q : measurement (data) Bayes’ rule:
  find the maximum a-posteriori probability (MAP)
  p(Q) is constant if p(L) is assumed constant -> maximum likelihood (ML)

## Slide 19
  Maximum Likelihood (ML)
  L = {lj} -> projections
    cij = sensitivity of detector i for activity j (incl. attenuation)
  assume Q = {qi} are samples of a Poisson distribution
  -> ->
  ->
  when -> qi! is constant

## Slide 20
Iterative solution: ML-EM algorithm
->
-> use iterative optimization : expectation-maximization (EM)
FBP
iterative reconstruction

## Slide 21
PROJECTIONS
REPROJECTION
COMPARE
image k
-> correction factor
X
IMAGE
image (k+1)
Sj cij lj old
qi
lj old
Si cij ( )
qi
Sj cij ljold
Si cij
1

## Slide 22
FBP
iterative reconstruction
(34 iterations)

## Slide 23: Outline
Data Modeling
Statistical Reconstruction
Heuristics

## Slide 24: Heuristics: Sum Game
=4
2=3
3=2
4=1

## Slide 25: Let Us Guess
3
2
3
2
4
3
2
1
0
0
0
0
4
3
2
1
6
4
-2
2
Update a guess based on data differences
Guess 1
Guess 0
Guess 2
Error
Error

## Slide 26: Same Game
=4
2=3
3=2
4=1

## Slide 27: Ratio instead of Difference
3
2
3
2
4
3
2
1
1
1
1
1
4.2
2.8
1.8
1.2
3
2
3/5
7/5
Update a guess based on data differences
Guess 1
Guess 0
Guess 2
Ratio
Ratio

## Slide 28: Weighted Sum
=4
2=3
3=2
4=1

## Slide 29: Simultaneous Correction
Correction Factor = (0.2*10/3+0.1*4/3)/(0.2+0.1)
4
3
2
1
1
1
1
1
4/3
10/3
Guess 1
Guess 0
Ratio
Ratio
Guess
Truth

## Slide 30
PROJECTIONS
REPROJECTION
COMPARE
image k
-> correction factor
X
IMAGE
image (k+1)
Sj cij lj old
qi
lj old
Si cij ( )
qi
Sj cij ljold
Si cij
1

## Slide 31: MAP: Regularized Reconstruction

## Slide 32: L0 Norm Replaced by L1 Norm
Measurement
Sparsifying
Objective Function consists of two terms:
The first is data fidelity
The second is prior information
= || ||2+

## Slide 33: Homework
Correction Factor = (0.2*10/3+0.1*4/3)/(0.2+0.1)
4
3
2
1
1
1
1
1
4/3
10/3
Guess 1
Guess 0
Ratio
Ratio
Guess
Truth
Please compute the correction factors for the rest three pixels, & correct Guess 0, & read my heuristic paper on iterative methods.

## Slide 34: Further Reading (04-Heuristic-SART.pdf)