## Slide 1
Ultrasound Imaging
Ge Wang
AI-based X-ray Imaging System (AXIS) Lab
Rensselaer Polytechnic Institute, Troy, New York, USA
November 12, 2021

## Slide 2: MI Schedule for F21
Tue | Topic | Fri | Topic
08/31 | Introduction | 09/03 | Neuron, Network & Backpropagation
09/07 | No Class | 09/10 | Hands-on 1: Python, Colab, & TensorF
09/14 | Architectures & Training | 09/19 | Hands-on 2: MNIST Classification
09/21 | Image Quality I | 09/24 | CT Physics
09/28 | CT Reconstruction | 10/01 | Deep CT Reconstruction
10/05 | Hands-on 3: CT Networks | 10/08 | More CT Contents
10/12 | MRI Physics | 10/15 | Spin-Echo
10/19 | K-Space Theorem | 10/22 | More MRI Contents
10/26 | Mid Exam | 10/29 | Deep MRI Reconstruction
11/02 | Hands-on 4: MRI Networks | 11/05 | Nuclear Physics
11/09 | Nuclear Imaging | 11/12 | US Imaging
11/16 | Optical Imaging | 11/19 | Multimodality Imaging
11/23 | Hands-on 5: Image Conversion | 11/26 | Thanksgiving
11/30 | Image Quality II | 12/03 | Stability, Interpretability, & Others
12/07 | Overview | 12/10 | Final Exam
Office Hour: Teaching Day 5-6pm: https://rensselaer.webex.com/meet/wangg6

## Slide 3: Outline
US Basics
– Wave Equation
– Transduction & Measurement
US & Tissue Interaction
– Huygens Principle
– Interaction Mechanisms
US Imaging Modes
– A, M, & B Modes
– Array Imaging
Doppler Imaging
– CW & PW
– Color
Image Quality
– Resolution
– Artifacts

## Slide 4: Acoustic Wave Eq Made Easy
Time=t: Volumetric air element has a constant density ρ under a pressure gradient (horizontal)
Time=t+Δt: Volumetric air element is moved (horizontally)
We need to describe this local dynamics

## Slide 5: Acoustic Wave Eq Made Easy
Let us think p as the pressure change relative to the equilibrium pressure
Note that an alternative form of the wave equation can be expressed in terms of pressure p, instead of displacement h or u

## Slide 6: General Treatment Involved Other Factors
https://www.win.tue.nl/~sjoerdr/papers/boek.pdf

## Slide 7: Maxwell's Equations
Gauss' Law for electricity
The total of the electric field over the surface of a volume equals the total charge inside
Gauss' Law for magnetism
No magnetic charge, i.e., no monopoles
Faraday's Law of magnetic induction
The left side is the total voltage change round a loop, and the right side is a varying magnetic field through the loop
Ampere's Law plus Maxwell's displacement current
The left side is the total magnetic force around a loop, and the right side is the current through the circuit, plus a varying electric field through the loop (displacement current)

## Slide 8: EM Wave Equation Made Easy
www.santarosa.edu/~lwillia2/42/WaveEquationDerivation.pdf

## Slide 9: Why B Orthogonal to E & X Axes?
(2) B components along X axis must be zero!
(1) E Fields identical on the parallel planes
(3) B curl must be due to components orthogonal to E & X axis
(0) Start with this simple case (Just consider the y component here without loss of generality)
Eq. (4)

## Slide 10
  Linear wave equation
    Assume small amplitude (small p) and nonviscous medium
    general solution
    Example of a solution (plane progressive wave) (x = direction of wave propagation)

## Slide 11: Physical Principles
What are ultrasonic waves?
elasticity and inertia of medium
-> displacements (compression and rarefaction) in the direction of wave motion -> longitudinal compression waves (> 20 kHz)

## Slide 12
Generation of Ultrasonic Waves
  piezo-electric crystal
    deforms under influence of electric field
    induces electric field under influence of deformation
  ultrasonic transducer generation :
piezo-electric crystal + electric field -> deformation
detection :
piezo-electric crystal + deformation -> electric field
compression wave
compression wave

## Slide 13: Outline
US Basics
– Wave Equation
– Transduction & Measurement
US & Tissue Interaction
– Huygens Principle
– Interaction Mechanisms
US Imaging Modes
– A, M, & B Modes
– Array Imaging
Doppler Imaging
– CW & PW
– Color
Image Quality
– Resolution
– Artifacts

## Slide 14
Wave propagation in inhomogeneous media
  Huygens' principle
  The principle that any point on a wave front of sound may be regarded as the source of secondary waves and that the surface that is tangent to the secondary waves can be used to determine the future position of the wave front.

## Slide 16
  Interference between waves
    for a large number of coherent point sources (i.e., with same frequency and phase) the complex interference pattern is called diffraction
      example: 2 (or more) waves can amplify (constructive interference) or decrease (destructive interference) each other -> positions with constructive and positions with destructive interference
P
coherent
d
O1
O2
constructive
destructive

## Slide 17
shape is related to the geometry of the acoustic source
circular planar source
near field
far field -> constructive interference
rectangular planar source

## Slide 18
  Specular reflection and refraction at smooth tissue boundaries
    Reflection
    Refraction
Ii
1
2
3
4
Ir
qi
qr
c1
c2
c1 = l1.f
c2 = l2.f
Ii
1
2
3
4
It
qi
qt
c1
c2
c1 = l1.f
c2 = l2.f
in practice: no perfect specularity
-> significant reflection in cone around qr
c1>c2

## Slide 19
Wave propagation in homogeneous media
  Specific acoustic impedance of a homogeneous medium
  p = acoustic pressure v = particle velocity
  r = mass density c = acoustic wave velocity = l.f = l/T
Substance c (ms-1) Z = r c (106kg/m2s)
Air (25°) 346 0.000410
Fat 1450 1.38
Water (25°) 1493 1.48
Soft tissue 1540 1.63
Liver 1550 1.64
Blood (37°) 1570 1.67
Bone 4000 3.8 to 7.4
Aluminum 6320 17.0
air -> tissue -> bone: r c Z= r c

## Slide 20: Acoustic Intensity

## Slide 21
    Snell’s law
    Direction qt changes
    Amplitude A (= p0) changes as well (Intensity I ~ p02)
      Transmission coefficient
      Reflection coefficient (R = T–1)
refracted
wave
reflected
wave

## Slide 22: Reflection & Transmission
Z2

## Slide 23
  if Z1 << Z2 or Z1 >> Z2 -> large reflection
  example:
      Zbone = rbone . cbone is large
      Zair = rair . cair ~ 0
  at tissue/bone transition -> large reflection (50%)
  at air/tissue transition -> total reflection (100%) -> use of gel

## Slide 24: Why Matching Layer?
The intensity transmitted can be computed as 1-R2 and
The intensity transmitted can be computed as T2?
Z2

## Slide 25
Acoustic intensity and loudness
  Acoustic intensity I = average energy (J) per time unit (s) through unit area (m2) perpendicular to propagation direction (W/m2)
  A tenfold increase in intensity I sounds a little more than twice as
  loud to the human ear
  -> to express the sound level L a logarithmic scale is used
  (unit of L is decibel (dB))
  I0 = the threshold (at 1000 Hz) of human hearing (absolute silence) ~ 0 dB
  example:
  intensity I x 10 ~ sound level L: + 10 dB (and perceived loudness: x 2)
  intensity I x 2 ~ sound level L: + 3 dB
  70-80 dB (vacuum cleaner) -> memory and concentration problems > 80 dB (heavy traffic (80dB) rock concert (120 dB) jumbo (140dB) -> damage > 120 dB -> pain threshold
10-12

## Slide 26
Attenuation
  Loss of energy during propagation because of viscosity -> heat
  a : attenuation coefficient
    in Np/cm (Np = Neper (dimensionless) calculated as ln(pz/p0))
    in dB/cm -> -> 1 Np = (20log10(e)) dB = 8.6859 dB
  a0 = a/f
    in Np/(cm.MHz) or dB/ /(cm.MHz)
  example: a0 = 0.5 dB/(cm.MHz) and f = 2 MHz -> amplitude/2 after 6 cm
Substance a0 (dB / cm.MHz)
Lung 41 (a is large for air and bone)
Bone 20
Kidney 1.0
Liver 0.94
Fat 0.63
Blood 0.18
Brain 0.85
Water 0.0022
n ~ 1 in medical imaging
z : propagation distance
f : frequency

## Slide 27
  Smallest possible inhomogeneity = point scatterer -> retransmits incident wave in all directions (Huygens)
  Finite scatterer = collection of point scatterers -> forms interference of constituting point scattering
  diffraction pattern diffraction pattern
  of small scatterer of large scatterer

## Slide 28
  Scattering
    homogeneous tissue inhomogeneous tissue
specular reflection
noise
scatter

## Slide 29: Outline
US Basics
– Wave Equation
– Transduction & Measurement
US & Tissue Interaction
– Huygens Principle
– Interaction Mechanisms
US Imaging Modes
– A, M, & B Modes
– Array Imaging
Doppler Imaging
– CW & PW
– Color
Image Quality
– Resolution
– Artifacts

## Slide 30: Gray scale Imaging
A mode (A = amplitude)
2 x distance = c x time
c for the different tissues

## Slide 31
M mode (M = motion)
  = repeated A mode -> gray scale image
distance
time

## Slide 32
B mode (B = brightness)

## Slide 33: Phased Arrays
Artificial phase shift after reception
-> dynamic focusing
-> simultaneous scan lines -> temporal resolution (70 - 80 Hz)
Linear array transducer
= 1D array of crystals
Phased array transducer
electronic sweeping
electronic focusing
Huygens'
principle

## Slide 34: Outline
US Basics
– Wave Equation
– Transduction & Measurement
US & Tissue Interaction
– Huygens Principle
– Interaction Mechanisms
US Imaging Modes
– A, M, & B Modes
– Array Imaging
Doppler Imaging
– CW & PW
– Color
Image Quality
– Resolution
– Artifacts

## Slide 37: Pulse Front & Scatter Meet at tis

## Slide 38: Pulse End (λ) & Scatter Meet at tie

## Slide 39: Echo Duration from Front to End

## Slide 40: Doppler Frequency fD

## Slide 41: Outline
US Basics
– Wave Equation
– Transduction & Measurement
US & Tissue Interaction
– Huygens Principle
– Interaction Mechanisms
US Imaging Modes
– A, M, & B Modes
– Array Imaging
Doppler Imaging
– CW & PW
– Color
Image Quality
– Resolution
– Artifacts

## Slide 42: Image Resolution

## Slide 43
Image reconstruction
  filtering
    removes high frequency noise
    removes low-freq. band in 2nd harmonic imaging
  envelope detection
    envelope ~ amplitude of the signal ~ gray value in the image

## Slide 44
  attenuation correction = time-gain compensation
    simple model = exponential decay
    different tissues ~ different attenuations -> most US scanners enable manual modification of the gain at different depths
  log-compression = gray scale transformation
    logarithmic function -> scatter (or speckle) is also visible
  scan conversion = sector reconstruction
    interpolation : polar grid -> rectangular grid

## Slide 45
Acquisition and reconstruction time
  example :
    depth : 20 cm -> travel distance = 40 cm
    c = 1540 m/s
    1 image = 120 scan lines
acquisition time = 32 ms (reconstruction in real time)
~ 30 images per second (30 Hz)
  Note :
    current scanners acquire multiple scan lines simultaneously temporal resolution : 70 - 80 Hz
è
è
è

## Slide 46
Artifacts
  interference pattern contains side lobes
  reverberations

## Slide 47: Homework
Optional: If we use two matching layers, will the transmission efficiency be further improved?
Read the Ultrasound imaging chapter through, especially the materials not fully covered in the class such as pages on image quality and artifacts